Content

- Scope

- Definitions and abbreviations

- Personnel qualifications

3.1 Medical fitness

3.2 Education and training

- Procedure

4.1 Principle

4.2 Samples

4.3 Equipment and materials

4.4 Reagents and solutions

4.5 Detailed procedure

4.6 Result and interpretation

4.7 Quality control

4.8 Waste management

- Related documents

|  | Compiled by | Examined by | Approved by | Replaced | New version |
| --- | --- | --- | --- | --- | --- |
| Name |  |  |  | Code: | Code: |
| Date |  |  |  |  |  |
| Signature |  |  |  |  |  |
| Laboratory area: |  | No of copies: | Reason for change: |  |  |

- Scope

This SOP describes the procedure for the biochemical test that detects the activity of catalase from bacterial colonies which is used for the identification of Mycobacterium tuberculosis.

Because the test involves handling suspensions containing high concentrations of viable, infectious bacteria, strict compliance with safety and protection measures is mandatory. The test must be performed in a laboratory meeting the WHO standards for biosafety level 3 with access restricted to authorized personnel only.

- Definitions and abbreviations

MW:	molecular weight

BSC:	biological safety cabinet

- Personnel qualifications

- Medical fitness

In accordance with national laws and practices, arrangements should be made for  appropriate health surveillance of TB laboratory workers:

- before enrolment in the TB laboratory;

- at regular intervals thereafter, annually or bi-annually;

- after any biohazard incident;

- at the onset of TB symptoms.

Ideally, individual medical records shall be kept for up to 10 years following the end of occupational exposure.

Laboratory workers should be educated about the symptoms of TB and provided with ready access to free medical care if symptoms arise.

Confidential HIV counselling and testing should be offered to laboratory workers. Options for reassignment of HIV-positive or immuno-suppressed individuals away from the high-risk areas of the TB laboratory should be considered.

All cases of disease or death identified in accordance with national laws and/or practice as resulting from occupational exposure to biological agents shall be notified to the competent authority.

- Education and training

Basic education and training must be given on the following topics:

- potential risks to health (symptoms of TB disease and transmission);

- precautions to be taken to minimize aerosol formation and prevent exposure;

- hygiene requirements;

- wearing and use of protective equipment and clothing;

- handling of potentially infectious materials;

- laboratory design, including airflow conditions;

- use of biological safety cabinets (operation, identification of malfunctions, maintenance);

- use of autoclaves, water-baths, pH meters (operation, identification of malfunctions, maintenance);

- prevention of incidents and steps to be taken by workers in the case of incidents (biohazard incidents, chemical, electrical and fire hazards);

- good laboratory practice and good microbiological techniques;

- organization of work flow;

- procedures;

- waste handling;

- importance of laboratory results for patient management;

- importance of laboratory results for the national TB programme.

The training shall be:

- given before a staff member takes up his/her post;

- strictly supervised;

- adapted to take account of new or changed conditions; and

- repeated periodically, preferably every year.

- Procedure

4.1 Principle

The 68 °C catalase test, which measures the heat stability of the enzyme activity, is one of the elements in the identification of M. tuberculosis and other tubercle bacilli.

Catalase converts hydrogen peroxide to water and oxygen, generating oxygen bubbles in a liquid solution. Virtually all mycobacteria except certain isoniazid-resistant tubercle bacilli are catalase-positive. Moreover, the catalase enzyme present in most of the mycobacterial species is heat-stable at pH 7.0  except in tubercle bacilli and a few other mycobacterial species such as M. fallax, M. chelonae, M. gastri, M. marinum, M. malmoense, M. haemophilum.

The test must be carried out on pure cultures otherwise it will yield false results.

4.2 Samples

Pure cultures of acid-fast bacilli, grown on solid media, minimum age 14 days

4.3 Equipment and materials

BSC, class I or II, annually certified

Water-bath with lid, heater and thermostat (or temperature block heater) at 68 °C

pH meter

Autoclave

Screw-cap reaction tubes, 13 x 85 mm

Reference strains of M. tuberculosis (or a previously identified strain) and a reference strain of M. terrae (saprophytic strain)

4.4 Reagents and solutions

4.4.1 Phosphate buffer, 0.067 mol/litre, pH 7.0

- Reagent 1. Dissolve 9.07 g KH2PO4 (monopotassium dihydrogen phosphate, MW 136.1) in 1000 ml distilled water.

- Reagent 2. Dissolve 23.68 g Na2HPO4.12H2O (disodium hydrogen phosphate dodecahydrate, MW 358.14)  in 1000 ml distilled water.

Mix 38.9 ml of reagent 1 with 61.1 ml of reagent 2. Check pH using a pH meter.

4.4.2 Hydrogen peroxide (H2O2), 30%

Not the 3% H2O2 obtainable from pharmacies.

Wear plastic gloves and a protective eye shield when handling H2O2.

4.4.3 Tween 80 solution, 10%

Mix 10 ml Tween 80 with 90 ml distilled water; autoclave at 121°C for 10 minutes. Store in the refrigerator up to 3 months.

4.4.4 Substrate

Mix equal volume of 30% hydrogen peroxide and 10% Tween 80. Always prepare the substrate fresh just before the test.

4.5 Detailed procedure

The entire procedure must be carried out in a biological safety cabinet.

- Dispense 0.5 ml of phosphate buffer, 0.067 ml/litre, pH 7.0, into two 16 x 125 mm screw-cap test-tubes.

- Emulsify several loopfuls of bacterial colonies scraped from solid media into the buffer.

- Place one tube in the water-bath at 68 °C for 20 minutes. Leave the other tube at room temperature.

- Allow the heated tube to cool to room temperature.

- Add 0.5 ml freshly prepared Tween 80–peroxide substrate to each tube and cap loosely.

- Observe for the formation of bubbles. Retain negative tubes for 20 minutes before discarding. Do not shake the tubes: Tween 80 alone may form bubbles on shaking, resulting in false-positive results.

4.6 Result and interpretation

- If oxygen bubbles form in the unheated tube but not in the heated tube, the test strain produces heat-labile catalase and probably belongs to the M. tuberculosis complex.

- If oxygen bubbles form in both the heated and unheated tubes, the test strain produces heat-stable catalase and is unlikely to be a member of the M. tuberculosis complex.

- If no oxygen bubbles form in either the heated or the unheated tube, either the test has failed or the strain may be catalase-negative (such as rare isoniazid-resistant strains of M. tuberculosis). Repeat the test with a larger quantity of bacilli to eliminate the possibility of a false-negative result.

4.7 Quality control

4.7.1 Morphology of colonies

Observe the morphology of colonies on solid medium to check culture purity.

4.7.2 Positive control: no bubble formation after heating

Carry out the test with a previously identified M. tuberculosis strain or a reference strain from a culture collection.

4.7.3 Negative control: bubble formation after heating.

Carry out the test with a reference strain such as M. terrae (saprophytic strain) from an international culture collection.

4.8 Waste management

Autoclave all contaminated materials.

- Related documents

Kent PT, Kubica GP. Public health mycobacteriology: a guide for the level III laboratory. Atlanta, GA, United States Department of Health and Human Services, Centers for Disease Control, 1985.

Laboratory services in tuberculosis control. Part III: Culture. Geneva, World Health Organization, 1998.

Rieder HL et al. Priorities for tuberculosis bacteriology services in low-income countries. Paris, International Union Against Tuberculosis and Lung Disease, 2007.

Vincent V, Gutiérrez MC. Mycobacterium: laboratory characteristics of slowly growing mycobacteria. In: Manual of clinical microbiology. Washington, DC, American Society for Microbiology, 2007:573–588.